FBD-MSP-ZSAV for chemiluminescence immunoassay

Examiner: Evaluation group in R&D department; 2016.11.16~12.16; at Suzhou Laboratory

- Examination systems

Equipment: AE-180 automated chemiluminescence immunoassay equipment

Biomarker: CEA, CA50 and HBSAg

The test group: FBD-MSP-ZSAV replacing MSP-SAV through OEM in issued reagent kits

The control group: the issued intact reagent kits themselves

Quantity: FBD-MSP-ZSAV at 10 g; MSP-SAV through OEM at 15g

Operation: following all those described in documents of the issued reagent kits

Range of references: calibrator S0 through S5, quality control sample, clinical specimens

Parameters: signal/noise ratios, sensitivity, linear ranges, reproducibility, consistence

- Results (original print screen result)

- Conclusion

3.1 The test group gave generally higher chemiluminescence signals than the control group, with some exceptionality. For CEA, somewhat lower signals were observed at higher levels.

3.2 Signal/noise ratios: basically consistent between the test group and the control group.

3.3 Reproducibility: the test group gave slightly larger CV than the control group.

3.4 Sensitivity: the test group gave comparable or higher sensitivity than the control group.

3.5 Linear range: consistent linearity within the tested ranges.

3.6 Consistence: one false positive result with HBsAg in clinical specimens.

3.7 Overall conclusion: FBD-MSP-ZSAV was applicable, after the adjustment of formulations.